HCal fADC Timing Resolution
3/13/2020
Scott Barcus

---
fADC Time Over Threshold
- Threshold crossing defines fADC time.
  - Threshold defined as ¼ average fADC peak height.
- fADC signals fit with Landau function.

---
fADC Reference Time
- Send two trigger copies over long cables to two fADC channels to get sloped edge for timing.
- Fit the leading edge of trigger copies with an exponential. (Linear fit was tested and worse.)
  - Ref time defined by threshold of fADC=1250.
  - PMT Time = fADC T.O.T. - ref time.

---
Timing Jitter Between Ref Ch.s
- Histogram of ref_time1 – ref_time2.
- Oddly shaped (other thresholds even worse).

---
Time Walk Effect
- Plot fADC signal integral vs. fADC T.O.T.
- Fit with exponential.

---
Time Walk Correction
- Corrected Time = fADC T.O.T. - ref time – f(exp fit)
- Both T.O.T. and ref time have distributions.
- σ2corrected time = σ2fADC T.O.T. + σ2ref time
- σref time not Gaussian.

---
Individual PMT Timing Resolutions
- Require vertical cosmics for best timing.
  - Three F1s fire vertically in a row and surrounding six F1s don't fire.

---
Time Walk Corrected Timing Resolutions
σfADC T.O.T. = (|σ2corrected time - σ2ref time|)1/2. (Gaus)

---
- Using histogram standard deviation.
  - 1.639 vs. 1.445 for Gaussian.

---
- Timing resolution worse lower down in HCal.
  - Probably just geometry. (histogram std. dev.)